SCPI Instrument Toolkit

Student User Manual

ESET 453 - Validation and Verification

Texas A&M University

April 21, 2026

# Table of Contents

Chapter 1: Getting Started
Chapter 2: Navigating the Official Docs
Chapter 3: Your First Measurement
Chapter 4: Finding and Running Examples
Chapter 5: Lab Report Workflow
Chapter 6: Variables, Calc, and Logging
Chapter 7: Scripting Mastery
Chapter 8: Multi-Instrument Workflows
Chapter 9: Data Analysis
Chapter 10: Instrument Tips and Gotchas
Chapter 11: Troubleshooting

# Chapter 1: Getting Started

## What is the SCPI Instrument Toolkit?

The SCPI Instrument Toolkit is a Python-based command-line tool for controlling lab instruments -- oscilloscopes, power supplies, multimeters, function generators, source measure units, and USB-to-I2C adapters -- directly from your terminal. Instead of clicking through vendor GUIs or writing full Python programs, you type short commands and get instant results.

The toolkit was built for ESET 453 (Validation and Verification) at Texas A&M University. It supports the instruments in the VOAL lab and works on both personal and managed TAMU machines.

## What You Need

Before installing, make sure you have:

- Python 3.10 or newer -- Check by running python --version in your terminal. - NI-VISA runtime -- Required for USB and GPIB communication with instruments. Download from https://www.ni.com/en/support/downloads/drivers/download.ni-visa.html

## Installation

### Standard install (any machine)

Open a terminal (PowerShell on Windows, Terminal on Mac/Linux) and run:

pip install git+https://github.com/T-O-M-Tool-Oauto-Mationator/scpi-instrument-toolkit.git

### TAMU managed machines (VOAL lab)

On TAMU-managed Windows machines where admin access is restricted, use this one-liner in PowerShell:

irm https://raw.githubusercontent.com/T-O-M-Tool-Oauto-Mationator/scpi-instrument-toolkit/main/setup-tamu.ps1 | iex

This handles PATH issues and installs the toolkit in your user Python directory.

### Verify installation

After installing, verify it works:

scpi-repl --version

You should see something like scpi-instrument-toolkit v1.0.17.

If scpi-repl is not recognized, use the module form instead:

python -m lab_instruments --version

## Launching the REPL

There are three ways to start the REPL:

With real instruments:

scpi-repl

This auto-discovers any instruments connected to your machine via USB, GPIB, or serial.

Without hardware (mock mode):

scpi-repl --mock

This simulates 14 instruments so you can practice commands without being in the lab.

Module form (if scpi-repl is not on your PATH):

python -m lab_instruments

python -m lab_instruments --mock

## What You See on First Launch

When you start the REPL, it automatically scans for connected instruments:

$ scpi-repl --mock

ESET Instrument REPL v1.0.17. Type 'help' for commands.

[INFO] Scanning for instruments in background...

[SUCCESS] Scan complete: found 14 device(s).

eset>

The eset> prompt means the REPL is ready for commands.

## Mock Mode: Practice Without Hardware

Mock mode simulates all the instruments you would find in the VOAL lab. This is useful for:

- Practicing commands before going to the lab - Writing and testing scripts at home - Learning the toolkit without hardware access

The 14 simulated devices are:

| Name | Type | Simulates |
| --- | --- | --- |
| psu1 | Power Supply | HP E3631A |
| psu2 | Power Supply | Matrix MPS-6010H |
| psu3 | Power Supply | Keysight EDU36311A |
| smu | Source Measure Unit | NI PXIe-4139 |
| awg1 | Function Generator | Keysight EDU33212A |
| awg2 | Function Generator | JDS6600 |
| awg3 | Function Generator | BK Precision 4063 |
| dmm1 | Multimeter | HP 34401A |
| dmm2 | Multimeter | OWON XDM1041 |
| dmm3 | Multimeter | Keysight EDU34450A |
| scope1 | Oscilloscope | Rigol DHO804 |
| scope2 | Oscilloscope | Tektronix MSO2024 |
| scope3 | Oscilloscope | Keysight DSOX1204G |
| ev2300 | USB-to-I2C Adapter | TI EV2300 |

Mock instruments return realistic random values (not constants), so your measurements will look like real data with small variations.

## Accessibility

If you use a screen reader or prefer plain text output without colors, use the --no-color flag:

scpi-repl --no-color

scpi-repl --mock --no-color

You can also set the NO_COLOR environment variable to disable colors globally.

## Connecting the STM32 Adapter (VOAL Lab Only)

The VOAL lab uses a small USB-to-I2C adapter to talk to the BQ76920 battery monitor board. TAMU-owned benches use a TI EV2300A; newer benches use an in-house STM32F405 replacement that plugs in the same way. Your REPL sees both as ev2300 and the commands are identical - you do not need to know which bridge is on your bench.

Assembled STM32 adapter (Feather + protoboard) with the JST cable that plugs into the BQ EVM

Plug-in order matters. Do this every time:

1. Set the bench PSU to 18 V and turn it on so the BQ EVM has power. 2. Plug the USB cable from the STM32 adapter (or EV2300) into your computer. 3. Press the BOOT button on the BQ EVM. This clears any stuck state on the I2C bus from the last student. 4. Launch scpi-repl and type scan. You should see ev2300 in the device list.

If scan does not find the adapter or reads return garbage, see Chapter 11 (Troubleshooting) - the fix is almost always "repeat the plug-in order above" or unplug-replug the USB cable.

You will never need to solder or disassemble the adapter. If a unit is physically broken, hand it to a TA - Chapter 12 of the TA Developer Guide covers the rebuild.

## Try It

Open a terminal and run:

scpi-repl --mock

Type list to see all connected mock instruments, then type help to see available commands. You are ready to start measuring.

# Chapter 2: Navigating the Official Docs

## The Docs Site

The official documentation lives at:

https://t-o-m-tool-oauto-mationator.github.io/scpi-instrument-toolkit/

This site is always up to date with the latest version of the toolkit. Bookmark it -- you will reference it often during labs.

## Site Structure

The docs are organized into these sections:

Home -- Quick start guide, installation, and overview of all commands.

Installation -- Detailed install instructions for Windows, macOS, and Linux, including TAMU managed machines.

Supported Instruments -- List of all supported models with connection details.

REPL section -- One page per instrument type, plus scripting, examples, and logging:

- General Commands -- scan, list, use, idn, raw, state, all off, sleep, exit - Power Supply (PSU) -- chan, set, meas, get, track, save/recall - Function Generator (AWG) -- chan, wave, freq, amp, offset, duty, phase - Multimeter (DMM) -- config, read, meas, check, beep, display - Oscilloscope (Scope) -- autoset, run/stop, chan, trigger, meas, waveform capture - Source Measure Unit (SMU) -- on/off, set, meas, sweep - EV2300 (USB-to-I2C) -- read_word, write_word, scan, probe - Scripting -- variables, loops, if/else, scripts, debugging - Example Workflows -- 15 ready-to-run examples with explanations - Log & Calc -- measurement log, calculations, CSV export - Plotting -- static plots, live plots, interactive charts

Python API -- How to control instruments from Python scripts (for advanced users).

Reference -- Auto-generated command lists and expression function tables.

Architecture -- How the toolkit is structured internally (for developers).

Troubleshooting -- Fixes for common errors and TAMU-specific issues.

## How to Search

The docs site has a search bar at the top. Type any keyword to find relevant pages:

- Type "voltage" to find PSU voltage commands, DMM voltage measurement, etc. - Type "frequency" to find AWG frequency settings, scope frequency measurement, etc. - Type "csv" to find how to export data to CSV files

The search is fast and searches across all pages.

## How to Find Instrument Commands

Each instrument type has its own page with every command documented. To find a command:

1. Go to the REPL section in the left sidebar 2. Click the instrument type (PSU, DMM, Scope, etc.) 3. Scroll to the command you need 4. Each command shows: syntax, parameters table, and usage examples

For example, to learn how to set a PSU voltage: - Go to "Power Supply (PSU)" page - Find the "psu set" section - You will see: psu set <channel> <voltage> [current] - Below that are examples: psu set 1 5.0, psu set 1 5.0 0.5

## The REPL Command Reference

Under the Reference section, there is an auto-generated "REPL Command Reference" page. This page lists every single command available in the REPL, organized by instrument type. It is generated directly from the source code, so it is always accurate and complete.

Use this page when you need a quick lookup of all commands for a specific instrument without reading the full documentation page.

## Expression Functions Reference

Also under Reference, the "Expression Functions" page lists every function you can use in calc expressions and variable assignments:

- Math: sqrt, sin, cos, tan, log, log10, exp, ceil, floor, abs, round - Type conversion: int, float, str, bool, hex, bin, oct - Constants: pi, e, inf, nan, True, False - Comparison: ==, !=, <, <=, >, >= - Boolean: and, or, not

## Using Help Inside the REPL

You do not always need to open the docs site. The REPL has built-in help:

help              # list all available commands

help psu          # show help for PSU commands

help dmm          # show help for DMM commands

help scope        # show help for scope commands

The docs command opens the full documentation in your browser:

docs              # opens the docs site in your default browser

## Finding and Running Examples

The examples command inside the REPL lists all 15 bundled examples:

examples          # list all available examples with descriptions

To load and run one:

examples load psu_dmm_test

script run psu_dmm_test

See Chapter 4 for detailed instructions on running examples.

## Try It

1. Open the docs site in your browser 2. Use the search bar to find "psu set" 3. Read the PSU page to understand the command syntax 4. Launch scpi-repl --mock and try the commands you found

# Chapter 3: Your First Measurement

This chapter walks through a complete measurement workflow step by step. You will set a PSU voltage, measure it with a DMM, log the result, and shut down safely.

You can follow along in mock mode (scpi-repl --mock) or with real instruments in the lab.

## Step 1: Launch the REPL

scpi-repl --mock

The REPL starts, scans for instruments, and shows the prompt:

eset>

## Step 2: See What is Connected

list

This prints all discovered instruments with their names and models:

<!-- doc-test: skip reason="sample list output, not runnable REPL input" -->

Name     Type               Model

psu1     Power Supply       HP E3631A

dmm1     Multimeter         HP 34401A

scope1   Oscilloscope       Rigol DHO804

...

The names (psu1, dmm1, scope1) are what you use in commands.

## Step 3: Select a Power Supply

use psu1

This sets psu1 as the "active" PSU. Now you can type psu commands without specifying which PSU:

[SUCCESS] Active: psu1

Alternatively, you can always prefix commands with the instrument name directly: psu1 set 1 5.0.

## Step 4: Enable the Output

psu chan 1 on

This enables channel 1 of the power supply. The output is now live (or simulated in mock mode).

[SUCCESS] Output enabled

## Step 5: Set the Voltage

psu set 1 5.0

This sets channel 1 to 5.0 volts. You can optionally set a current limit:

psu set 1 5.0 0.5

This sets 5.0 V with a 0.5 A current limit.

[SUCCESS] Set 1: 5.0V

## Step 6: Measure the PSU Output

psu meas 1 v

This reads the actual output voltage from the PSU's internal measurement circuitry:

5.001753 V

Note: the value is slightly different from 5.0 -- this is normal. Real instruments have small measurement variations.

## Step 7: Configure the DMM

dmm1 config vdc

This configures the DMM (dmm1) for DC voltage measurement. The mode stays active until you change it.

[SUCCESS] Configured for dc_voltage

## Step 8: Read the DMM and Store the Result

voltage = dmm1 meas unit=V

This does three things: 1. Takes a DC voltage reading from the DMM 2. Stores the result in a variable called voltage 3. Records it in the measurement log with the label "voltage" and unit "V"

Output:

voltage = 5.000241 V

The variable voltage is now available for calculations and print statements.

## Step 9: Print the Result

print "Measured voltage: {voltage} V"

Output:

Measured voltage: 5.000241 V

The {voltage} placeholder is replaced with the stored value.

## Step 10: View the Measurement Log

log print

This shows all recorded measurements in a table:

Label      Value       Unit   Source

voltage    5.000241    V      dmm1

You can export this to a CSV file:

log save results.csv

## Step 11: Shut Down Safely

all off

This disables all instrument outputs at once -- a safe shutdown.

exit

This closes the REPL and disconnects from all instruments.

## Putting It All Together

Here is the complete sequence you just ran:

scpi-repl --mock

list

use psu1

psu chan 1 on

psu set 1 5.0

psu meas 1 v

dmm1 config vdc

voltage = dmm1 meas unit=V

print "Measured voltage: {voltage} V"

log print

all off

exit

## Common Mistakes

Forgetting to enable the output. If you set a voltage but do not run psu chan 1 on, the output stays at 0 V. Always enable the channel before setting voltage.

Forgetting to configure the DMM mode. If you run dmm1 meas without first running dmm1 config vdc, the DMM may be in the wrong mode (resistance, frequency, etc.) and return unexpected values.

**Using psu meas v instead of the assignment syntax.** Plain psu meas 1 v prints the value but does NOT save it. Use label = psu1 meas v unit=V to record it in the log.

Not specifying the channel number. psu set 5.0 is ambiguous. Always specify the channel: psu set 1 5.0.

## Try It

Run the complete sequence above in mock mode. Then try:

1. Change the voltage to 3.3 V and measure again 2. Measure the current: current = psu1 meas i unit=A 3. Calculate power: calc power = voltage * current unit=W 4. View the updated log: log print

# Chapter 4: Finding and Running Examples

The toolkit ships with 15 bundled example scripts that demonstrate common lab workflows. You can run them as-is, customize them with parameters, or use them as starting points for your own scripts.

## Listing Available Examples

examples

This prints all 15 examples with their descriptions:

psu_dmm_test         Set PSU to a voltage, measure with DMM, log result

voltage_sweep        Sweep PSU through a list of voltages, log DMM reading at each step

awg_scope_check      Output a sine wave on AWG, measure frequency and PK2PK on scope

freq_sweep           Sweep AWG through frequencies, scope measures each

psu_ramp             Ramp PSU voltage from start to end in steps

live_voltage_sweep   Voltage sweep with live plot

live_multi_plot      Multi-channel live plot

live_freq_sweep      Frequency sweep with live plot

live_combined_plot   Combined live plot with PSU and DMM

cross_script_demo    Cross-script demo calling Python from SCPI

conditional_psu_check  PSU check with if/elif/else branching

assert_limits        Assert and check with pass/fail limits

while_counter        While loop with counter and accumulator

syntax_reference     Full syntax tour of all scripting features

complete_cross_script  Complete cross-script demo

## Loading an Example

examples load psu_dmm_test

This copies the example script into your session. You can now see its source:

script show psu_dmm_test

Output shows every line of the script with syntax highlighting.

To load all examples at once:

examples load all

## Running an Example

script run psu_dmm_test

This runs the script from top to bottom. In mock mode, you will see:

=== PSU/DMM Voltage Test ===

Target: 5.0V

[SUCCESS] Output enabled

[SUCCESS] Set 1: 5.0V

[INFO] Sleeping 0.5s...

psu_v = 5.001753 V

[SUCCESS] Configured for dc_voltage

vtest = 5.000241 V

=== Test complete ===

Label      Value       Unit   Source

psu_v      5.001753    V      psu1

vtest      5.000241    V      dmm1

## Using Parameter Overrides

Many examples define default variables at the top (like voltage = 5.0). You can override these when running:

script run psu_dmm_test voltage=3.3 label=v3v3

This sets voltage to 3.3 and label to "v3v3" before running the script. The log entry will use your custom label:

Label      Value       Unit   Source

psu_v      3.298176    V      psu1

v3v3       3.300122    V      dmm1

## Walkthrough: psu_dmm_test

This is the simplest example. Here is what each line does:

voltage = 5.0                    # default voltage (overridable)

label = vtest                    # default label (overridable)

print "=== PSU/DMM Voltage Test ==="

print "Target: {voltage}V"       # prints with variable substitution

psu1 chan 1 on                   # enable PSU output

psu1 set 1 {voltage}            # set PSU to target voltage

sleep 0.5                        # wait 500ms for output to settle

psu_v = psu1 meas v unit=V      # measure PSU's own output

dmm1 config vdc                  # configure DMM for DC voltage

{label} = dmm1 meas unit=V      # measure with DMM, use custom label

print "=== Test complete ==="

log print                        # show all measurements

After running, you can compute the error between PSU setpoint and DMM reading:

<!-- doc-test: skip reason="depends on the vtest label recorded by the preceding script" -->

calc error = {vtest} - 5.0 unit=V

calc error_pct = ({vtest} - 5.0) / 5.0 * 100 unit=%

## Walkthrough: voltage_sweep

This example sweeps through multiple voltages and records the DMM reading at each step:

examples load voltage_sweep

script run voltage_sweep

What it does: 1. Turns on PSU channel 1 2. Configures the DMM for DC voltage 3. For each voltage in the list (1.0, 2.0, 3.3, 5.0, 9.0, 12.0): - Sets the PSU to that voltage - Waits 500ms for settling - Records the DMM reading as v_1.0, v_2.0, etc. 4. Turns off the PSU 5. Prints the log and saves to CSV

The log will have entries like:

Label      Value       Unit   Source

v_1.0      0.999847    V      dmm1

v_2.0      1.999623    V      dmm1

v_3.3      3.299814    V      dmm1

v_5.0      4.999952    V      dmm1

v_9.0      8.999431    V      dmm1

v_12.0     11.999684   V      dmm1

## Modifying an Example

To customize an example, edit it after loading:

examples load psu_dmm_test

script edit psu_dmm_test

This opens the script in your system text editor. Make changes, save, and close. Then run:

script run psu_dmm_test

## Saving Custom Scripts

If you modify an example and want to keep it:

script save

This writes all in-memory scripts to the scripts directory as .scpi files. They will be automatically loaded next time you start the REPL.

## Try It

1. Launch mock mode: scpi-repl --mock 2. Load and run the psu_dmm_test: examples load psu_dmm_test then script run psu_dmm_test 3. Run it again with different parameters: script run psu_dmm_test voltage=3.3 label=v3v3 4. Load the voltage_sweep: examples load voltage_sweep then script run voltage_sweep 5. View the CSV output: log save sweep_results.csv 6. Try the syntax_reference example to see every scripting feature: examples load syntax_reference then script run syntax_reference

# Chapter 5: Lab Report Workflow

This chapter walks through the complete workflow from connecting instruments to producing figures for your lab report.

## The Workflow

1. Connect instruments and launch the REPL 2. Configure measurement modes 3. Take measurements and log them 4. Export data to CSV 5. Load CSV in Python with pandas 6. Plot with matplotlib 7. Include figures in your DOCX report

## Step 1: Connect and Launch

Plug in your instruments via USB. Open a terminal and run:

scpi-repl

The REPL scans and finds your instruments automatically. Run list to confirm.

## Step 2: Configure Instruments

Set up each instrument for your test:

use psu1

psu chan 1 on

psu set 1 5.0 0.5

dmm1 config vdc

## Step 3: Take Measurements

Use the assignment syntax to record every reading with a label and unit:

v_set = psu1 meas v unit=V

i_set = psu1 meas i unit=A

v_dmm = dmm1 meas unit=V

For a sweep, use a for loop:

for v 1.0 2.0 3.3 5.0

psu set 1 {v}

sleep 0.5

v_{v} = dmm1 meas unit=V

end

## Step 4: Calculate Derived Values

Use calc to compute quantities from your measurements:

<!-- doc-test: skip reason="depends on v_dmm and i_set recorded by preceding measurement steps" -->

calc power = v_dmm * i_set unit=W

calc error = v_dmm - 5.0 unit=V

calc error_pct = (v_dmm - 5.0) / 5.0 * 100 unit=%

## Step 5: Review and Export

View all measurements:

log print

Export to CSV:

log save lab3_data.csv

The CSV file contains columns: Label, Value, Unit, Source. You can open it in Excel or Python.

## Step 6: Plot in Python

Create a Python script (e.g., plot_lab3.py) to generate figures:

import pandas as pd

import matplotlib.pyplot as plt

# Load the CSV

df = pd.read_csv("lab3_data.csv")

# Filter voltage sweep data

sweep = df[df["Label"].str.startswith("v_")]

voltages = [float(label.split("_")[1]) for label in sweep["Label"]]

measured = sweep["Value"].values

# Plot

plt.figure(figsize=(8, 5))

plt.plot(voltages, measured, "bo-", label="Measured")

plt.plot(voltages, voltages, "r--", label="Ideal")

plt.xlabel("Set Voltage (V)")

plt.ylabel("Measured Voltage (V)")

plt.title("PSU Voltage Accuracy")

plt.legend()

plt.grid(True)

plt.tight_layout()

plt.savefig("voltage_accuracy.png", dpi=150)

plt.show()

Run the script:

python plot_lab3.py

This generates voltage_accuracy.png -- a publication-ready figure for your report.

## Step 7: Include in Your Report

Insert the PNG figure into your DOCX lab report. Add a caption referencing the test conditions:

<!-- doc-test: skip reason="prose caption example, not REPL input" -->

"Figure 1: PSU voltage accuracy measured with HP 34401A DMM.

Set voltage vs. measured voltage across 1.0-5.0 V range."

## Complete Example Script

Here is a full REPL script you can save and reuse:

# lab3_voltage_test.scpi

log clear

psu1 chan 1 on

dmm1 config vdc

for v 1.0 2.0 3.3 5.0

psu1 set 1 {v}

sleep 0.5

v_{v} = dmm1 meas unit=V

end

psu1 chan 1 off

log print

log save lab3_data.csv

print "Data saved to lab3_data.csv"

## Tips for Good Lab Data

- Always wait for the PSU output to settle before measuring. Use sleep 0.5 (500ms) minimum. - Record both the PSU's own measurement and the DMM's measurement. The difference tells you about measurement accuracy. - Use descriptive labels: v_3v3_rail is better than v1. - Run log clear at the start of each test to avoid mixing data from different runs. - Save CSV files with descriptive names including the date: lab3_voltage_sweep_20260416.csv.

## Try It

1. Run the complete example script above in mock mode 2. Export to CSV with log save 3. Write a Python script to plot the results 4. Generate a PNG figure

# Chapter 6: Variables, Calc, and Logging

## Variables

Variables store values for reuse in commands, calculations, and print statements.

### Assigning Variables

voltage = 5.0

current = 0.25

label = "vtest"

count = 0

Variables can hold numbers (int or float), strings, booleans, and lists.

### Using Variables

Use {varname} to insert a variable's value into a command or print statement:

psu set 1 {voltage}

print "Setting {voltage} V"

### Arithmetic in Assignments

You can do math directly in assignments:

<!-- doc-test: skip reason="reference example -- voltage/current aren't set in this block" -->

power = voltage * current

doubled = voltage * 2

half = voltage / 2

remainder = 17 % 5

squared = voltage ** 2

### Available Operators

| Operator | Meaning | Example | Result |
| --- | --- | --- | --- |
| + | Addition | 3 + 4 | 7 |
| - | Subtraction | 10 - 3 | 7 |
| * | Multiplication | 6 * 7 | 42 |
| / | Division | 15 / 4 | 3.75 |
| // | Floor division | 15 // 4 | 3 |
| % | Modulo | 17 % 5 | 2 |
| ** | Power | 2 ** 10 | 1024 |

### Increment and Decrement

count = 0

count++          # count is now 1

count++          # count is now 2

count--          # count is now 1

### Compound Assignment

x = 10

x += 5           # x is now 15

x -= 3           # x is now 12

x *= 2           # x is now 24

x /= 4           # x is now 6

### Math Functions

You can use math functions in any expression:

sq = sqrt(144)           # 12.0

angle = degrees(pi)      # 180.0

lg = log10(1000)         # 3.0

rounded = round(3.14159, 2)  # 3.14

Full list: sqrt, sin, cos, tan, asin, acos, atan, log, log2, log10, exp, ceil, floor, abs, round, pow, min, max, sum, degrees, radians, hypot.

Constants: pi, e, inf, nan, True, False.

### Type Conversions

x = int(3.7)         # 3

y = float(5)         # 5.0

h = hex(255)         # "0xff"

b = bin(10)          # "0b1010"

flag = bool(1)       # True

## The Measurement Log

Every time you use the assignment syntax with an instrument, the result is saved to the measurement log:

v_out = psu1 meas v unit=V       # saved to log as "v_out"

dmm_reading = dmm1 meas unit=V   # saved to log as "dmm_reading"

The log is a table with columns: Label, Value, Unit, Source.

### Viewing the Log

log print

Output:

Label          Value       Unit   Source

v_out          5.001753    V      psu1

dmm_reading    5.000241    V      dmm1

### Clearing the Log

log clear

Removes all entries. Use this before starting a new test.

### Exporting the Log

log save results.csv          # CSV format (opens in Excel)

log save results.txt          # formatted text table

## Calc: Derived Calculations

The calc command computes a value from existing measurements and stores the result in both the log and as a variable.

### Basic Calc

v = psu1 meas v unit=V

i = psu1 meas i unit=A

calc power = v * i unit=W

The result appears in log print alongside your measurements.

### Referencing Variables by Name

Calc can reference any variable or measurement label by its bare name:

voltage = 5.0

current = 0.25

calc power = voltage * current unit=W

### The "last" Keyword

dmm1 config vdc

reading = dmm1 meas unit=V

calc doubled = last * 2 unit=V

last always refers to the most recently stored measurement value.

### Chained Calculations

Results from calc are stored as variables, so you can chain them:

v_in = psu1 meas v unit=V

i_in = psu1 meas i unit=A

v_out = dmm1 meas unit=V

calc power_in = v_in * i_in unit=W

calc power_out = v_out * i_in unit=W

calc efficiency = power_out / power_in * 100 unit=%

### Common Calc Patterns

Percentage error:

measured = dmm1 meas unit=V

calc error_pct = (measured - 5.0) / 5.0 * 100 unit=%

Voltage gain (dB):

v_in = scope1 meas 1 PK2PK unit=V

v_out = scope1 meas 2 PK2PK unit=V

calc gain = v_out / v_in

calc gain_db = 20 * log10(gain) unit=dB

Resistance from V and I:

v = psu1 meas v unit=V

dmm1 config idc

i = dmm1 meas unit=A

calc resistance = v / i unit=ohms

## Print Command

Print text with variable substitution:

print "Voltage is {voltage} V"

print "Power = {power} W, Efficiency = {efficiency}%"

Quotes are optional but recommended for clarity.

## Try It

1. Launch mock mode and set some variables 2. Do arithmetic: power = 5.0 * 0.25 3. Measure from PSU and DMM, store with labels 4. Compute error percentage with calc 5. View everything with log print 6. Export with log save test.csv

# Chapter 7: Scripting Mastery

Scripts let you automate entire test sequences. Instead of typing commands one at a time, you write them in a file and run them with one command.

## Creating Scripts

### Method 1: Script Editor

script new my_test

This opens your system text editor. Write commands, one per line. Save and close.

### Method 2: Record from Interactive Session

record start my_test

psu1 chan 1 on

psu1 set 1 5.0

sleep 0.5

v = psu1 meas v unit=V

log print

record stop

Every command you type between record start and record stop is saved to the script.

### Method 3: Load from File

script run ./my_test.scpi

Run a .scpi file directly from disk without copying it to the scripts library.

## Running Scripts

script run my_test                    # run by name

script run my_test voltage=3.3       # with parameter override

script run ./lab3.scpi                # run by file path

## Variables in Scripts

Define default values at the top of your script. Users can override them with key=value on the command line.

# my_test.scpi

voltage = 5.0

label = vtest

psu1 chan 1 on

psu1 set 1 {voltage}

sleep 0.5

{label} = dmm1 meas unit=V

log print

Run with defaults:

script run my_test

Run with overrides:

script run my_test voltage=3.3 label=v3v3

## For Loops

Repeat a block of commands for each value in a list:

for v 1.0 2.0 3.3 5.0

psu1 set 1 {v}

sleep 0.5

v_{v} = dmm1 meas unit=V

end

The loop variable {v} takes each value in order: 1.0, then 2.0, then 3.3, then 5.0.

### Linspace for Smooth Sweeps

Generate evenly spaced values:

RAMP = linspace 0 5.0 11

This creates 11 values from 0 to 5.0: 0 0.5 1.0 1.5 2.0 2.5 3.0 3.5 4.0 4.5 5.0

Use it in a for loop:

RAMP = linspace 0 5.0 11

for v RAMP

psu1 set 1 {v}

sleep 0.3

v_{v} = dmm1 meas unit=V

end

### Multi-Variable For Loops

Iterate over paired values:

for ch,v 1,5.0 2,12.0 3,3.3

psu1 set {ch} {v}

end

## While Loops

Repeat until a condition is false:

i = 0

total = 0

while i < 5

i++

total += i

end

print "Sum 1..5 = {total}"

### While with Break and Continue

j = 0

while j < 20

j++

if j == 3

continue       # skip iteration 3

end

if j > 6

break          # stop after 6

end

print "j = {j}"

end

## If / Elif / Else

Conditional branching:

voltage = 5.05

if voltage > 5.1

verdict = "OVER"

elif voltage < 4.9

verdict = "UNDER"

else

verdict = "OK"

end

print "Voltage {voltage} V: {verdict}"

## Assert (Hard Stop)

Assert stops the script immediately if the condition is false:

<!-- doc-test: skip reason="reference example -- voltage not set in this block" -->

assert voltage > 0 "voltage must be positive"

assert voltage < 6.0 "voltage below safety limit"

Use assert for safety-critical checks that must pass before continuing.

## Check (Soft Test)

Check records a PASS or FAIL result but does NOT stop the script:

<!-- doc-test: skip reason="reference example -- voltage not set in this block" -->

check voltage > 4.9 "above lower bound"

check voltage < 5.1 "below upper bound"

View results with:

log report

Check is useful for validation testing where you want to run all tests and report results at the end.

## Sleep and Pause

<!-- doc-test: skip reason="pause prompts operator for interactive Enter" -->

sleep 0.5              # wait 500ms (for instrument settling)

sleep 2                # wait 2 seconds

pause                  # wait for user to press Enter

pause "Connect DUT"    # wait with custom message

## Repeat

Repeat a block a fixed number of times:

repeat 5

v = dmm1 meas unit=V

sleep 1

end

## Calling Sub-Scripts

<!-- doc-test: skip reason="requires a saved sub-script named other_script on disk" -->

call other_script

Runs other_script from within the current script. Variables are shared between caller and callee.

## The Debugger

Step through a script one command at a time:

script debug my_test

Debugger commands:

| Command | Action |
| --- | --- |
| n / next | Execute next line |
| c / cont | Continue to next breakpoint |
| s / step | Same as next |
| p / print | Show all variables |
| q / quit | Stop debugging |

### Breakpoints

Add a breakpoint in your script:

<!-- doc-test: skip reason="breakpoint only fires under script debug, not at the onecmd prompt" -->

psu1 set 1 5.0

breakpoint                # debugger stops here

v = dmm1 meas unit=V

When running with script debug, execution pauses at each breakpoint so you can inspect variables.

## Error Handling

set -e            # stop script on any error

set +e            # continue despite errors (default)

With set -e, the script stops at the first error (failed measurement, invalid command, etc.). Without it, errors are printed but the script continues.

## Example: Complete Test Script

# lab3_full_test.scpi

set -e

log clear

voltage = 5.0

tolerance = 5

print "=== Lab 3: Voltage Accuracy Test ==="

print "Target: {voltage} V, Tolerance: {tolerance}%"

psu1 chan 1 on

psu1 set 1 {voltage}

sleep 1.0

dmm1 config vdc

measured = dmm1 meas unit=V

calc error_pct = (measured - {voltage}) / {voltage} * 100 unit=%

check error_pct -5 5 "within {tolerance}% tolerance"

psu1 chan 1 off

log print

log save lab3_results.csv

log report

print "=== Test Complete ==="

## Try It

1. Create a script with script new my_sweep 2. Write a for loop that sweeps voltage and measures 3. Run it in mock mode 4. Try the debugger: script debug my_sweep 5. Add a breakpoint and inspect variables

# Chapter 8: Multi-Instrument Workflows

Real lab tests often involve multiple instruments working together. This chapter shows common multi-instrument patterns.

## Addressing Multiple Instruments

When you have multiple instruments of the same type, they are numbered:

list

# psu1: HP E3631A

# psu2: Matrix MPS-6010H

# dmm1: HP 34401A

# dmm2: OWON XDM1041

# scope1: Rigol DHO804

Address any instrument directly by prefixing its name:

<!-- doc-test: skip reason="reference example -- requires a multi-PSU / multi-DMM lab setup (psu2, dmm2)" -->

psu1 set 1 5.0

psu2 set 1 12.0

dmm1 config vdc

dmm2 config idc

Or use use to switch the active instrument for short commands:

use psu1

psu set 1 5.0       # acts on psu1

use psu2

psu set 1 12.0      # acts on psu2

## Pattern 1: PSU + DMM Voltage Verification

Verify that the PSU is outputting the correct voltage using an independent DMM measurement.

# Set the PSU

psu1 chan 1 on

psu1 set 1 5.0

sleep 0.5

# Measure with both instruments

psu_v = psu1 meas v unit=V

dmm1 config vdc

dmm_v = dmm1 meas unit=V

# Compare

calc error = psu_v - dmm_v unit=V

calc error_pct = (psu_v - dmm_v) / psu_v * 100 unit=%

print "PSU reads {psu_v} V, DMM reads {dmm_v} V"

print "Error: {error_pct}%"

log print

## Pattern 2: AWG + Scope Signal Check

Generate a waveform with the AWG and verify it on the scope.

# Configure AWG

awg1 chan 1 on

awg1 wave 1 sine freq=1000 amp=2.0 offset=0

sleep 0.5

# Configure and capture on scope

scope1 autoset

sleep 1.0

# Measure signal characteristics

freq = scope1 meas 1 FREQUENCY unit=Hz

pk2pk = scope1 meas 1 PK2PK unit=V

# Verify

calc freq_error = (freq - 1000) / 1000 * 100 unit=%

calc amp_error = (pk2pk - 2.0) / 2.0 * 100 unit=%

print "Frequency: {freq} Hz (error: {freq_error}%)"

print "Amplitude: {pk2pk} Vpp (error: {amp_error}%)"

# Cleanup

awg1 chan 1 off

log print

## Pattern 3: PSU Voltage Sweep with DMM Readback

Sweep the PSU through a range of voltages and record DMM readings at each step.

log clear

psu1 chan 1 on

dmm1 config vdc

for v 1.0 2.0 3.3 5.0 9.0 12.0

psu1 set 1 {v}

sleep 0.5

psu_v = psu1 meas v unit=V

dmm_v = dmm1 meas unit=V

calc error_{v} = psu_v - dmm_v unit=V

end

psu1 chan 1 off

log print

log save voltage_sweep.csv

## Pattern 4: AWG Frequency Sweep with Scope Measurement

Characterize frequency response by sweeping the AWG and measuring on the scope.

awg1 chan 1 on

awg1 wave 1 sine amp=2.0 offset=0

for f 100 500 1000 5000 10000 50000

awg1 freq 1 {f}

sleep 0.5

freq_{f} = scope1 meas 1 FREQUENCY unit=Hz

pk2pk_{f} = scope1 meas 1 PK2PK unit=V

end

awg1 chan 1 off

log print

log save freq_response.csv

## Pattern 5: Power Measurement (PSU + DMM)

Measure both voltage and current to compute power dissipation.

psu1 chan 1 on

psu1 set 1 5.0

# Use PSU for voltage, DMM for current

v = psu1 meas v unit=V

dmm1 config idc

i = dmm1 meas unit=A

calc power = v * i unit=W

calc resistance = v / i unit=ohms

print "V = {v} V, I = {i} A"

print "P = {power} W, R = {resistance} ohms"

log print

## Pattern 6: Dual-Supply Configuration

Set up positive and negative rails for an op-amp circuit.

<!-- doc-test: skip reason="requires dual-PSU lab setup; psu1 track + psu2 not present in single-PSU mock" -->

# Positive rail

psu1 chan 1 on

psu1 set 1 12.0

# Negative rail (if using tracking mode)

psu1 track on          # channels mirror: +12V and -12V

# Or use two separate PSUs

psu1 chan 1 on

psu1 set 1 12.0        # +12V rail

psu2 chan 1 on

psu2 set 1 12.0        # -12V rail (wire negative)

# Verify both rails

pos_v = psu1 meas v unit=V

neg_v = psu2 meas v unit=V

print "Positive rail: {pos_v} V"

print "Negative rail: {neg_v} V"

## Pattern 7: Full V&V Test Sequence

A complete validation test that uses multiple instruments:

<!-- doc-test: skip reason="requires 2x DMM lab setup and a real DUT at 3.3 V (check fails against unconfigured mock)" -->

set -e

log clear

print "=== V&V Test: DUT Power-Up Sequence ==="

# Step 1: Power up DUT

psu1 chan 1 on

psu1 set 1 3.3 0.5

sleep 1.0

# Step 2: Verify supply voltage

dmm1 config vdc

v_supply = dmm1 meas unit=V

check v_supply 3.135 3.465 "supply within 5% of 3.3V"

# Step 3: Check quiescent current

dmm2 config idc

i_quiescent = dmm2 meas unit=A

check i_quiescent 0 0.050 "quiescent current < 50mA"

# Step 4: Apply test signal

awg1 chan 1 on

awg1 wave 1 sine freq=1000 amp=1.0 offset=1.65

sleep 0.5

# Step 5: Verify output

scope1 autoset

sleep 1.0

out_freq = scope1 meas 1 FREQUENCY unit=Hz

out_pk2pk = scope1 meas 2 PK2PK unit=V

# Step 6: Cleanup and report

awg1 chan 1 off

psu1 chan 1 off

log print

log report

log save vv_test_results.csv

print "=== Test Complete ==="

## Safe Shutdown

Always shut down instruments at the end of a test:

all off          # disables all outputs on all instruments

Or shut down individually:

<!-- doc-test: skip reason="references psu2 which isn't in the single-PSU mock" -->

psu1 chan 1 off

awg1 off

psu2 chan 1 off

## Try It

1. In mock mode, run Pattern 1 (PSU + DMM verification) 2. Run Pattern 2 (AWG + scope signal check) 3. Try the full V&V test sequence 4. Export results and review the CSV

# Chapter 9: Data Analysis

After collecting measurements, you need to analyze the data and produce figures for your lab report. This chapter shows how to work with exported CSV data in Python.

## Exporting Data

From the REPL, export your measurement log:

log save lab_data.csv

This creates a CSV file with columns: Label, Value, Unit, Source.

## Loading CSV with pandas

import pandas as pd

df = pd.read_csv("lab_data.csv")

print(df)

Output:

Label           Value       Unit    Source

v_1.0           0.999847    V       dmm1

v_2.0           1.999623    V       dmm1

v_3.3           3.299814    V       dmm1

v_5.0           4.999952    V       dmm1

## Basic Statistics

print(f"Mean: {df['Value'].mean():.4f}")

print(f"Std:  {df['Value'].std():.4f}")

print(f"Min:  {df['Value'].min():.4f}")

print(f"Max:  {df['Value'].max():.4f}")

## Filtering Data

Select specific measurements by label pattern:

# All voltage sweep points

sweep = df[df["Label"].str.startswith("v_")]

# Only measurements from dmm1

dmm_data = df[df["Source"] == "dmm1"]

# Only voltage measurements

volts = df[df["Unit"] == "V"]

## Plotting with matplotlib

### Line Plot

import matplotlib.pyplot as plt

voltages_set = [1.0, 2.0, 3.3, 5.0]

voltages_measured = sweep["Value"].values

plt.figure(figsize=(8, 5))

plt.plot(voltages_set, voltages_measured, "bo-", label="Measured")

plt.plot(voltages_set, voltages_set, "r--", label="Ideal")

plt.xlabel("Set Voltage (V)")

plt.ylabel("Measured Voltage (V)")

plt.title("PSU Voltage Accuracy")

plt.legend()

plt.grid(True)

plt.tight_layout()

plt.savefig("voltage_accuracy.png", dpi=150)

### Scatter Plot

plt.figure(figsize=(8, 5))

plt.scatter(voltages_set, voltages_measured, c="blue", s=50)

plt.xlabel("Set Voltage (V)")

plt.ylabel("Measured Voltage (V)")

plt.title("Voltage Correlation")

plt.grid(True)

plt.tight_layout()

plt.savefig("voltage_scatter.png", dpi=150)

### Bar Chart (Error Comparison)

errors = [m - s for m, s in zip(voltages_measured, voltages_set)]

error_pct = [(m - s) / s * 100 for m, s in zip(voltages_measured, voltages_set)]

plt.figure(figsize=(8, 5))

plt.bar([str(v) for v in voltages_set], error_pct, color="steelblue")

plt.xlabel("Set Voltage (V)")

plt.ylabel("Error (%)")

plt.title("Percentage Error by Voltage")

plt.axhline(y=0, color="black", linewidth=0.5)

plt.grid(axis="y")

plt.tight_layout()

plt.savefig("error_bar.png", dpi=150)

### Frequency Response (Bode-like)

freqs = [100, 500, 1000, 5000, 10000, 50000]

gains_db = [0.1, 0.0, -0.2, -1.5, -3.1, -8.2]  # example data

plt.figure(figsize=(8, 5))

plt.semilogx(freqs, gains_db, "bo-")

plt.xlabel("Frequency (Hz)")

plt.ylabel("Gain (dB)")

plt.title("Frequency Response")

plt.grid(True, which="both")

plt.axhline(y=-3, color="red", linestyle="--", label="-3 dB")

plt.legend()

plt.tight_layout()

plt.savefig("freq_response.png", dpi=150)

## Computing Error Metrics

import numpy as np

set_values = np.array(voltages_set)

measured = np.array(voltages_measured)

# Absolute error

abs_error = np.abs(measured - set_values)

print(f"Max absolute error: {abs_error.max():.6f} V")

# Percentage error

pct_error = np.abs((measured - set_values) / set_values) * 100

print(f"Max percentage error: {pct_error.max():.4f}%")

# RMS error

rms_error = np.sqrt(np.mean((measured - set_values) ** 2))

print(f"RMS error: {rms_error:.6f} V")

## Combining Multiple Test Runs

If you ran the same test multiple times:

run1 = pd.read_csv("run1.csv")

run2 = pd.read_csv("run2.csv")

run3 = pd.read_csv("run3.csv")

# Combine

all_runs = pd.concat([run1, run2, run3], keys=["Run 1", "Run 2", "Run 3"])

# Group by label and compute statistics

grouped = all_runs.groupby("Label")["Value"]

summary = grouped.agg(["mean", "std", "min", "max"])

print(summary)

## Saving Figures for Reports

Always use these settings for professional-looking figures:

plt.savefig("figure_name.png", dpi=150, bbox_inches="tight")

- Use dpi=150 for print quality (300 for high-res) - Use bbox_inches="tight" to remove excess whitespace - Use descriptive filenames: lab3_voltage_accuracy.png

## Template: Complete Analysis Script

import pandas as pd

import matplotlib.pyplot as plt

import numpy as np

# Load data

df = pd.read_csv("lab_data.csv")

# Extract sweep data

sweep = df[df["Label"].str.startswith("v_")]

set_v = [float(label.split("_")[1]) for label in sweep["Label"]]

meas_v = sweep["Value"].values

# Compute errors

abs_err = np.abs(meas_v - set_v)

pct_err = np.abs((meas_v - set_v) / set_v) * 100

# Print summary

print("=== Voltage Accuracy Summary ===")

for s, m, e in zip(set_v, meas_v, pct_err):

print(f"  {s:5.1f} V -> {m:.6f} V  (error: {e:.4f}%)")

print(f"  Max error: {pct_err.max():.4f}%")

print(f"  RMS error: {np.sqrt(np.mean(abs_err**2)):.6f} V")

# Plot

fig, (ax1, ax2) = plt.subplots(1, 2, figsize=(12, 5))

ax1.plot(set_v, meas_v, "bo-", label="Measured")

ax1.plot(set_v, set_v, "r--", label="Ideal")

ax1.set_xlabel("Set Voltage (V)")

ax1.set_ylabel("Measured Voltage (V)")

ax1.set_title("Voltage Accuracy")

ax1.legend()

ax1.grid(True)

ax2.bar([str(v) for v in set_v], pct_err, color="steelblue")

ax2.set_xlabel("Set Voltage (V)")

ax2.set_ylabel("Error (%)")

ax2.set_title("Percentage Error")

ax2.grid(axis="y")

plt.tight_layout()

plt.savefig("lab_report_figure.png", dpi=150)

plt.show()

## Try It

1. Run a voltage sweep in mock mode and export to CSV 2. Copy the template analysis script above 3. Run it against your CSV data 4. Generate figures for a report

# Chapter 10: Instrument Tips and Gotchas

Each instrument has quirks. This chapter covers the things that trip students up most often.

## Power Supplies

### HP E3631A (psu1)

- Triple output with shared ground. Channels 1 (0-6V), 2 (0-25V), and 3 (0 to -25V) share a common ground. You cannot float them independently. - Tracking mode links channels 2 and 3 for +/- supply configurations. Use psu track on. - No current readback below ~1mA. Very low currents read as 0. Use an external DMM for milliamp-level measurements. - Output enable is per-channel. You must run psu chan 1 on before psu set 1 5.0 produces any output.

### Keysight EDU36311A (psu3)

- OVP/OCP protection. Over-voltage and over-current protection can trip and latch off the output. Check the front panel if the output unexpectedly drops to 0. - 5 save/recall slots. Use psu save 1 and psu recall 1 to store and restore setups. - USB enumeration delay. After connecting, wait 5-10 seconds before running scan. The USB interface takes time to initialize.

### Matrix MPS-6010H (psu2)

- Serial interface with slow command response. Allow 100-200ms between commands. The toolkit handles this automatically, but manual raw commands may need sleep 0.2 between them. - Remote mode must be enabled. The toolkit does this automatically on first command.

## Multimeters

### HP 34401A (dmm1)

- NPLC setting matters. Integration time (power line cycles) directly affects accuracy and speed: - nplc=0.02 -- fast but noisy (50 readings/second) - nplc=1 -- good balance (default) - nplc=10 -- high accuracy, slow (about 0.5 seconds per reading) - nplc=100 -- maximum accuracy, very slow (about 2 seconds per reading) - Display blanking. dmm display off speeds up measurements by ~10% on GPIB. - Auto-range settling. When auto-ranging, the first reading after a range change may be inaccurate. Take a throwaway reading first. - GPIB address. Default is usually address 22. Check the front panel under Utility > Remote Interface.

### Keysight EDU34450A (dmm3)

- No NPLC support. Uses SLOW/MEDIUM/FAST speed settings instead. - Capacitance and temperature modes. Supported but not on all HP models -- check before using. - Dual display. Can show two measurements simultaneously on the front panel, but the REPL reads only the primary display.

### OWON XDM1041 (dmm2)

- 4.5 digits. Lower resolution than the HP 34401A (6.5 digits). Expect more noise in readings. - Limited modes. Does not support NPLC, display text, or fetch.

## Oscilloscopes

### Rigol DHO804 (scope1)

- Autoset works well but may choose unexpected time/voltage scales for complex signals. Always verify the display after autoset. - Measurement force. After scope single + scope wait_stop, run scope meas_force before reading measurements. This forces the DSP to finalize values. - 9.9e+37 sentinel. If a measurement returns 9.9e+37, the scope has not triggered yet. Use scope wait_stop to wait for the trigger. - Built-in AWG. The DHO804 has a built-in function generator accessible through scope1 awg_* commands. - Counter and DVM. Built-in frequency counter and digital voltmeter -- useful for quick checks without switching instruments.

### Tektronix MSO2024 (scope2)

- Basic command set only. Does not support counter, DVM, cursors, recording, or built-in AWG features. - Slower autoset. Allow 2-3 seconds for autoset to complete. - save_waveform_csv accepts lists. You can pass [1, 2] to save multiple channels in one CSV.

### Keysight DSOX1204G (scope3)

- Most feature-rich scope. Supports counter, DVM, cursors, mask test, built-in AWG (WGEN), screenshot, display, acquire, math, labels, segmented acquisition, and measurement statistics. - No recording support. Cannot use the record/playback feature. - Segmented memory. Supports capturing multiple trigger events in separate memory segments.

## Function Generators

### Keysight EDU33212A (awg1)

- Dual channel. Both channels can output independently. Use awg chan 1 on and awg chan 2 on. - PRBS mode. Supports pseudo-random binary sequence for BER testing. Other AWGs do not have this. - High impedance vs 50 ohm. The AWG assumes a 50 ohm load. If your circuit is high-impedance, the actual output voltage will be double the set amplitude.

### JDS6600 (awg2)

- Serial interface. Slower than USB instruments. Allow 200ms between commands. - 17 waveform types. Supports more waveform shapes than other AWGs. - Frequency accuracy. DDS-based, so frequency accuracy is excellent.

### BK Precision 4063 (awg3)

- Single channel only. No channel 2.

## Source Measure Unit

### NI PXIe-4139 (smu)

- PXIe boot order matters. The PXIe chassis must be powered on before the host PC boots. If the PC boots first, the PXIe card will not be detected. Power cycle the chassis and reboot. - 4-quadrant operation. Can source and sink both voltage and current. Be careful with polarity. - Compliance limits. Always set current compliance when sourcing voltage, and voltage compliance when sourcing current. The toolkit enforces this. - nidcpower required. The NI-DCPower Python driver must be installed. See the NI PXIe-4139 setup page in the docs.

## EV2300 (USB-to-I2C Adapter)

### TI EV2300

- HID interface. Communicates over USB HID, not VISA. No NI-VISA required. - I2C address format. Addresses are 7-bit. Use hex: ev2300 read_word 0x08 0x09. - Bus recovery. If I2C communication hangs, use ev2300 fix to attempt bus recovery. - Slow for bulk transfers. HID has limited bandwidth. Block reads of large registers take time.

## General Tips

- **Always sleep after setting a voltage. PSU outputs need 100-500ms to settle. Without this delay, your DMM will read an intermediate value. - Configure the DMM mode before measuring. The most common mistake is measuring without setting the correct mode (vdc, idc, res, etc.). - Use all off before disconnecting. This safely disables all outputs before you unplug anything. - Save early, save often.** Run log save frequently. If the REPL crashes or you accidentally run log clear, your data is gone. - Use descriptive labels. ch1_vout_5v0 is much better than v1 when you look at the CSV later.

# Chapter 11: Troubleshooting

## "scpi-repl" is not recognized

Windows (personal machine): Run python -m lab_instruments once. The toolkit automatically adds the Scripts folder to your PATH. Open a new terminal window and try again.

TAMU managed machines (VOAL): Registry edits are blocked. Use the module form as your permanent launch method:

python -m lab_instruments

python -m lab_instruments --mock

Or set PATH for the current session only:

$env:PATH = "$env:LOCALAPPDATA\Programs\Python\Python312\Scripts;" + $env:PATH

Then scpi-repl will work in that terminal window.

## "No module named 'lab_instruments'"

The toolkit is not installed. Run:

pip install git+https://github.com/T-O-M-Tool-Oauto-Mationator/scpi-instrument-toolkit.git

If you get permission errors, add --user:

pip install --user git+https://github.com/T-O-M-Tool-Oauto-Mationator/scpi-instrument-toolkit.git

## "Scan complete: no instruments found"

- Check that instruments are powered on and connected via USB/GPIB. - Make sure NI-VISA is installed (required for VISA instruments). - Check that no other program (BQStudio, vendor GUI, another REPL session) is using the instrument. Only one program can connect at a time. - Try force_scan -- this does a slower, more thorough scan. - On Linux, make sure your user is in the dialout group for serial instruments.

## "Resource busy" or "Cannot open resource"

Another program is already connected to the instrument. Close it first:

- Another terminal running scpi-repl - A Python script using the same instrument - Vendor software (BQStudio, Keysight Connection Expert, NI MAX, etc.) - A previous REPL session that crashed without disconnecting

If a previous session crashed:

# On macOS/Linux -- find and kill stuck processes

ps aux | grep scpi

kill <pid>

# On Windows

tasklist | findstr python

taskkill /PID <pid> /F

## Measurements return "9.9e+37"

This is the SCPI sentinel value for "measurement not available." Common causes:

- The scope has not triggered yet. Use scope wait_stop before reading measurements. - The DMM is not configured. Run dmm config vdc (or appropriate mode) first. - The instrument is in an error state. Run idn to verify communication, then state reset.

## PSU output stays at 0 V

- Did you enable the output? Run psu chan 1 on first. - Is OVP/OCP tripped? Check the front panel for protection indicators. - Did you set the correct channel? psu set 1 5.0 sets channel 1. psu set 2 5.0 sets channel 2.

## DMM reads the wrong value

- Check the measurement mode. Run dmm config vdc for DC voltage, dmm config idc for DC current, etc. - Wrong range? Auto-range may pick an unexpected range. Specify manually: dmm config vdc 10. - Probe connected to wrong terminals? Voltage uses V/COM terminals. Current uses A/COM or mA/COM.

## "NI-VISA not found"

Install the NI-VISA runtime:

https://www.ni.com/en/support/downloads/drivers/download.ni-visa.html

On Linux, install from the .deb or .rpm package. On macOS, install from the .dmg.

After installing, restart your terminal.

## NI PXIe-4139 not detected

1. Power cycle the PXIe chassis FIRST 2. Then reboot the host PC 3. The chassis must be on before the PC boots -- boot order matters 4. Verify with NI MAX that the device appears

## Serial port permission denied (Linux)

Add your user to the dialout group:

sudo usermod -a -G dialout $USER

Log out and back in for the change to take effect.

## EV2300 communication errors

The I2C bus may be stuck. Try:

ev2300 fix

This sends clock pulses to recover the bus. If that does not work, disconnect and reconnect the EV2300 USB cable.

## "git" is not recognized (Windows)

Install Git for Windows from https://git-scm.com/download/win or use the toolkit's setup script which does not require git.

## REPL crashes or freezes

- Check if an instrument disconnected during a measurement. Reconnect and run scan. - If the REPL is frozen, press Ctrl+C to interrupt, then Ctrl+D to exit. - If Ctrl+C does not work, close the terminal window.

## Update errors

If pip install --upgrade fails:

pip install --force-reinstall git+https://github.com/T-O-M-Tool-Oauto-Mationator/scpi-instrument-toolkit.git

## Getting Help

- Official docs: https://t-o-m-tool-oauto-mationator.github.io/scpi-instrument-toolkit/ - In the REPL: type help for command list, help psu for specific help - Report bugs: https://github.com/T-O-M-Tool-Oauto-Mationator/scpi-instrument-toolkit/issues